## Workings (hidden)
Fridges and Freezers
| Age |  |  | Mechanically Cooled? |  |  | Alarms? |  |  | Ambient temp>21C |  |  | Filters Clean? |  |  | Doors Opened |  |  | Operating temperature |  |  |  | How Frosted?
|  | ULT | -20 C/fridge |  | ULT | -20 C |  | ULT | -20 C |  | ULT | -20 C/fridge |  | ULT | -20 C |  | ULT | -20 C/fridge |  | ULT |  | -20 C |  | ULT | -20 C/fridge
| (select) |  |  | (select) |  |  | (select) |  |  | (select) |  |  | (select) |  |  | (select) |  |  | (select) |  | (select) |  | (select)
| 0 - 2 years | 10.5 | 1.5 | Yes | 75 | 5 | Yes | 70 | 70 | > 25 C | 0.15 | 0.15 | Yes | 0 | 0 | Infrequently | 0 | 0 | -85 C | 0.14 | -25 C | 0.15 | Frosted | 0.08 | 0.08
| 2 - 6 years | 12 | 2 | No | 0 | 0 | No | 0 | 0 | 21 - 25 C | 0.1 | 0.1 | No | 0.12 | 0.12 | Moderately | 0.075 | 0.06 | -80 C | 0 | -22 C | 0.08 | Seals Defrosted | 0.05 | 0.05
| > 6 years | 15 | 3 | [6 empty] | < 21 C | 0 | 0 |  |  |  | Frequently | 0.15 | 0.12 | -75 C | -0.14 | -20 C | 0 | Defrosted | 0 | 0
[19 empty] | -70 C | -0.28
[19 empty] | -60 C | -0.38
CO2 Incubators | [9 empty] | Safety Cabinets |  |  |  | Computers
| Approx size (L) | Associated kWh/day | Age | change | Doors opened? |  | Operating Temperature |  |  |  | Ducted? |  |  | Type of computer |  | number of monitors |  | Screen brightness |  | Time until sleep |  | Usage type
| (select) |  | (select) |  | (select) |  | (select) |  |  |  | (select) |  |  | (select) |  | (select) |  | (select) |  | (select) |  | (select)
| <60 | 1.1 | New | 1 | Infrequently | 0 | 37 C | 0 |  |  | Yes | 1.25 |  | Laptop | 0.07 | 1 | 0 | Low | 0.6 | 1 min | 0.016667 | Basic | 1
| 120-180 | 1.9 | 1-5 years | 1.5 | Moderately | 0.075 | 38 - 50 C | 0.12 |  |  | No | 1 |  | Desktop PC | 0.13 | 2 | 0.06 | Mid | 0.8 | 5 mins | 0.083333 | Advanced | 1.75
| 200-250 | 2.7 | > 5 years | 1.85 | Frequently | 0.15 | > 50 C | 0.2 | [7 empty] | 3 | 0.12 | Bright | 1 | 10 mins | 0.16667
[20 empty] | 15 mins | 0.25
Fan motor consumption |  |  | FC |  | Region | % heating | % cooling | [12 empty] | 20 mins | 0.33333
Extract motor size (kW) | 0.55 | standard KW for 1500 FC | Make-up air? |  | (select) |  |  |  |  | Waste | Bin bag mass options | Assumed Average mass |  | Research Area |  | Funding | Lab type |  |  | 30 mins | 0.5
Number of extract motors | 1 |  | (select) |  | Scotland | 0.7 | 0.1 |  |  |  | (select) |  |  | (select) |  | (select) | (select) |  |  | 1 hr | 1
Assumed motor efficiency | 75% |  | Yes | 1.25 | North-East | 0.7 | 0.1 |  |  |  | n/a | 0 |  | Chemistry |  | Yes | Clinical Trials |  |  | 2 hrs | 2
|  |  | No | 1 | North-West | 0.7 | 0.1 |  |  |  | 0 - 3 kg | 2 |  | Computing |  | No | Core Facility |  |  | never
Make-up motor size (kW) | 0.37 |  |  |  | Yorkshire | 0.7 | 0.1 |  |  |  | 4 - 6 kg | 5 |  | Engineering |  |  | Private Research
Number of make-up motors | 1 |  |  |  | Northern Ireland | 0.7 | 0.1 |  |  |  | 7 - 9 kg | 8 |  | Life Sciences |  |  | Public Research
Assumed motor efficiency | 80% |  |  |  | East Midlands | 0.6 | 0.2 |  |  |  | 10 - 12 kg | 11 |  | Medical Sciences |  |  | Teaching
[5 empty] | West Midlands | 0.6 | 0.2 | [6 empty] | Physics
Heating/Cooling cost |  |  |  |  | Wales | 0.6 | 0.2
Average dT in heating season | 10 | assume 10°C |  |  | East of England | 0.6 | 0.2
Boiler efficiency, assume | 80% |  |  |  | South-East | 0.5 | 0.3
Average dT in cooling season | 3 | assume 3°C |  |  | South-West | 0.5 | 0.3
Chiller COP, assume | 3 |  |  |  | London | 0.5 | 0.3
Specific heat capacity of air | 1.2 | (kWh/m3/oC)

## 1. Assumptions
Details to be entered by LEAF participants | [6 empty] | Details to be entered by LEAF administrators
|  | Laboratory Context |  |  |  |  | The fields below are used to calculate costs and carbon emissions.
These values are defaults used at UCL, and are correct as of August 2019.
Number of Full Time staff within the lab | [6 empty] | You can update these for your institution if values are different.
Space within the lab (m2) | [6 empty] | Electricity cost | 0.13 | per kWh
Total time spent implementing LEAF (months) | [6 empty] | Gas cost | 0.03 | per kWh
[7 empty] | DEFRA CO2 factor for Grid Electricity | 0.2556 | kg CO2e /kWh
| Wellcome | UKRI | Charities | Institutional | Private | Other
Receive funding from: | (select) | (select) | (select) | (select) | (select) | (select) | DEFRA CO2 factor for Natural Gas | 0.18385 | kg CO2e /kWh
Research area of your laboratory: | (select) | [5 empty] | ULT Freezer Maintenance | 120 | per unit
Type of laboratory: | (select) | If teaching lab, for how many months do you operate? |  |  |  | 12
[7 empty] | -20 Freezer/Fridge Maintenance | 35 | per unit
[7 empty] | Freezer/Fridge Alarms | 70 | per unit
[7 empty] | Water Cost (total Supply & Sewerage) | 2 | per m3
Waste Data
[8 empty] | Cost per Tonne | kg Carbon per Tonne
[7 empty] | General Waste | 118.25 | 21.8
[7 empty] | Recycling | 107.25 | 21.8
[7 empty] | Clinical Waste | 495 | 21.8
[7 empty] | Offensive Waste | 148.5 | 21.8
Note on E9: If you don’t know the precise floor space, estimate.
Note on E11: Varying this will not affect final savings estimates, but rather is for record keeping.
Note on H18: If your lab is not in use all year (e.g. teaching labs), please indicate the number of months it is active (1-12 months). Your lab must be listed as a "teaching" lab in the left hand drop-down bar for this to apply to savings.

## 3. Silver
Silver
Total Criteria: 16
Category | # | Accreditation Level | Criteria | How did you meet this? | Useful Links
Waste | 17 | Silver | The lab has assessed its use of consumables and implemented realistic measures to reduce use. These efforts should target single-use plastics where feasible. | Write how you've addressed this here, or why it's not applicable |  | P
Waste | 18 | Silver | There is a minimum contamination of recycling in clinical waste bins (no more than 10%), and lab members are aware of best practice.
People | 19 | Silver | The lab has communicated with other groups/labs/departments about sustainable practices, and/or has taken part in a sustainability audit.
Purchasing | 20 | Silver | The lab is aware and makes use of schemes offered by suppliers/manufacturers which increase reuse, recycling, and waste reduction. This includes but is not limited to tip box recycling and the return of polystyrene boxes and Winchesters to suppliers.
Equipment | 21 | Silver | Freezers, fridges, and LN2 dewars are maintained or there is a plan in place going forward to achieve this. This includes defrosting, removing unwanted samples, checking seals, and cleaning filters on ULT freezers.
Equipment | 22 | Silver | Washers, autoclaves, and any equipment which permits batching, are only run when full. The lab considers appropriate sizing when buying such equipment.
Equipment | 23 | Silver | There is a system in place permitting the booking and sharing of communal equipment.
Equipment | 24 | Silver | Where feasible, freezers and fridges have temperatures raised and drying cabinets/ovens have had temperatures lowered.
IT | 25 | Silver | Critical data has been backed up, or there is a plan moving forward to ensure it is.
Sample & Chemical Management | 26 | Silver | Procedures for equipment breakdown are in place and well communicated to minimise losses. This may include but is not limited to freezer alarms, back-up storage spaces identified, call-out procedures, etc..
Sample & Chemical Management | 27 | Silver | The 12 Principles of Green Chemistry have been considered for current lab members, and communicated to the new members.
Research Quality * | 28 | Silver | The lab is aware of any relevant local core facilities or equivalents. Either there is a valid rationale for not utilising such a facility, or the lab makes regular use of them.
Research Quality * | 29 | Silver | The lab has a forum for sharing and discussing negative results.
Ventilation | 30 | Silver | Fume cupboards and Local Exhaust Ventilation equipment is not used for extended storage, and nothing impedes internal airflow.
Ventilation | 31 | Silver | Users have been made aware and have improved sash lowering of fume cupboards, and/or turn safety cabinets off when finished (at least 80% of the time).
Water | 32 | Silver | Sustainable water use is communicated to all lab users. This includes specifying what levels of water purity are necessary for various applications and ways to avoid taps running (e.g. soaking glassware).
Teaching * | 28 | Silver | An awareness of resource use and associated environmental impacts is incorporated into practical laboratory learning and teaching.
Note on E4: Criteria Below you will find all criteria you should address for Silver. If you scroll your mouse over each criteria, background on why this criteria is important will appear as a comment.
Note on F4: How? Please write in how your lab has addressed each of the criteria below. You will find tips and advice as comments when you hover your mouse over each box.
Note on G4: Useful Links Click on links below which may facilitate implementing the criteria, including free resources.
Note on E5: Why? Many common lab methods and processes rely on single use plastics. Much of the plastics have been treated rendering them difficult to recycle, or may come into contact with contaminated materials. As a result of this immense amounts of plastic waste are produced. Much of this plastic must be incinerated, directly harming our atmosphere.
Note on F5: Tips & Advice: • Consider glass as a possible alternative to some plastic items. • Opt for packaging which minimises polystyrene and wrapping. • In terms of priority reducing plastic usage should be the primary target, with reuse secondary, then recycling, and finally incineration when necessary.
Note on E6: Why? Mixing of waste streams can result both in a H&S hazard (through contamination of non-hazardous streams) or result in waste of energy (through excess treatment of non-hazardous waste). Ensuring users are aware of how to separate their waste is a crucial step to ensure contamination is minimised.
Note on F6: Tips & Advice: • Ensure all people who work in the lab receive training on using appropriate bins. • Check if there is existing signage available, or create your own clear and simple signage. • Check if there have been any waste audits conducted, or conduct your own audits with lab members to highlight if waste streams are being correctly utilised.
Note on E7: Why? One of the best ways to spread good practice is simply by word of mouth and local engagement.
Note on F7: Tips & Advice: • Reach out to local sustainability departments to see if you can assist with peer-audits of other labs. • Estimate impact of your efforts, and share results with colleagues and social media. • Share with senior management benefits of improving your lab's sustainability and efficiency.
Note on E8: Why? Often suppliers and manufacturers will offer schemes which facilitate recycling, reuse, or reductions in packaging. Users are encouraged to engage with their suppliers on such options, and take advantage. Should suppliers not offer such schemes, consider alternative suppliers or pressure them to provide these services.
Note on F8: Tips & Advice: • Schemes which facilitate reduce/reuse should be prioritised over recycling schemes. E.g. tip-box recycling schemes are extremely similar to general recycling. Focus on polystyrene take-back, or carbon off-setting schemes. • Check with stores/facilities it there are existing schemes which you may adopt. • Some schemes may require a certain volume of materials to be collected; engage with your neighbouring labs to collect a big enough volume.
Note on E9: Why? Cold storage uses a huge amount of energy. Each ultra-low temperature freezer can consume more energy than the average household! Well maintained equipment will use less energy, last longer, and crucially keep samples safer.
Note on F9: Tips & Advice: • Seal defrosts should be prioritised over entire defrosts to ensure internal temperatures are maintained, but units should never be so frosted over that the contents are inaccessible. • Some service contracts may cover parts or all of this criteria. Consider if you have or desire this type of contract. • If not covered through a service contract, such duties should be shared around or captured via a job description. Consider a rota.
Note on E10: Why? Autoclaves and washers can be extremely energy and water intensive, and thus should only be operated when necessary. Ensuring you have the correctly sized equipment can also reduce lag time if properly planned.
Note on F10: Tips & Advice: • Ask lab members if they believe there may be any future variation in usage of such equipment. • Set scheduled run times, and ensure users are aware of how to make best use.
Note on E11: Why? Booking systems have many benefits, though they primarily permit efficient equipment use. They also allow equipment usage to be tracked and monitored, facilitating maintenance and upkeep. A digital booking system can ensure all users have appropriate training prior to use.
Note on F11: Tips & Advice: • Check if there are existing booking systems in your facility/department. • Ensure the booking system allows for only appropriately trained lab members to book on such equipment. • Consider making equipment bookable to outside your laboratory/area.
Note on E12: Why? Equipment efficiencies will greatly decrease as they are operated at more extreme temperatures (very hot or cold). Particularly ULT freezers, which will use 30% less energy when at -70C vs. -80C. Each ULT freezer can use as much energy as the average house!
Note on F12: Tips & Advice: • Consider which equipment to focus on. Typically these are ovens which may be too hot, or freezers running too cold. Most standard freezers don't need to be colder than -20°C, and fridges should be 4°C. • With ULT freezers, be considerate of others. Engage prior to varying any temperature. Note that -70°C used to be the standard temperature. • It may be that -75°C is a middle ground. Freeze-thaw cycles can truly damage samples, so keep open door times to a minimum. Consider operating back-up units at warmer temperatures (to also deter misuse).
Note on E13: Why? Losses of data can represent immense amounts of laboratory research time, including the usage of resources. Avoiding a loss of such data can not only reduce reproduction of assays, but protect the career of a researcher.
Note on F13: Tips & Advice: • Consider the importance of lab-book like data. • Ensure your data back-up plans are aligned with any central IT directives. • Minimise the amount necessary to be backed-up to reduce server requirements.
Note on E14: Why? Loss of samples and chemicals can represent invaluable losses of data or funds. Ensure there are appropriate measures in place for realistic possible failures and minimise the risk of such losses.
Note on F14: Tips & Advice: • Consider what may happen if any member of staff finds for example a failed freezer. Post procedures where necessary. • Identify nearby back-up units with sufficient capacity, and ensure they are kept empty. Communal back-ups may suffice. • Consider remote alarm systems for high-value materials. Ensure there is an accompanying response strategy.
Note on E15: Why? The 12 Principles of Green Chemistry were created to drive more sustainable chemical reactions, products, and processes. It is up to users how to apply these, but staff should be reminded that there may be efficient alternatives.
Note on F15: Tips & Advice: • Post the 12 principles somewhere in the lab. • Raise the 12 principles at a lab meeting or forum where they may be discussed. • Check if others labs running similar experiments have already discovered more efficient methods.
Note on B16: This criteria is specifically for research labs. For teaching labs, the research quality criteria are not necessarily applicable.
Note on E16: Why? Core facilities can have a variety of names (e.g. technical platforms). Typically they should provide quicker turnover of samples, cheaper/efficient processes, and crucially ensure more consistency in processing. Such facilities will also contain experts in their field, maximising the potential outputs.
Note on F16: Tips & Advice: • Check your institute's local facilities for what services are available, or consult neighbour labs as to which facilities they use. Consider facilities outside your institute for high through-put. • Engage with a facility on sample preparation and methods. Agree what is required, and ensure your samples are correctly prepared. • Ensure core facilities are attributed for their input on any publications.
Note on B17: This criteria is specifically for research labs. For teaching labs, the research quality criteria are not necessarily applicable.
Note on E17: Why? "Negative results" is data which does not necessarily support an intended hypothesis. Such negative results, while not ideal for immediate publication, may reveal crucial information which could guide or negate the need for other experiments. The aim of this criteria is to allow an environment where negative results may be shared and discussed in a manner that facilitates better future research.
Note on F17: Tips & Advice: • A forum for negative results should record negative results for future lab members to consult. • The lab should consider making negative results available for external groups where possible, as this is likely to have the greatest impact. • Consider mandating that all research-talks must include 2 minutes on what went wrong.
Note on E18: Why? Materials left inside units may impede the flow of air. This can not only increase the risk of contamination, it can make the fan work harder than necessary wasting energy.
Note on F18: Tips & Advice: • If a fume cupboard is intended to store equipment, ensure that airflow testing is done when equipment is in place, and consider having containment testing done. • Check if there are alternative methods to store chemicals with suitable ventilation aside from a fume cupboard. Materials stored for extended periods may become a risk of explosion.
Note on E19: Why? Lowering the sash of a VAV fume cupboard and turning off a safety cabinet can result in significant savings, as well as being good practice in terms of cleanliness and H&S.
Note on F19: Tips & Advice: • Note some older fume cupboards may require users to push a button to engage lower fan speeds. Ensure your lab manager knows how your unit functions. • Consider a booking system to assist in reminding people when it's OK to turn off a safety cabinet. • Ensure reminders are positive - As frustrating as it may be that some users may forget, aim to lead through example and positivity. • Lowering a sash, even if it doesn't save energy, should always be encouraged.
Note on E20: Why? Users may be unsure what levels of purity are required, and as a result, tend to use excessively pure water for simple applications. Equally running taps can be avoided through training and allowing for soaking of items (e.g. glassware, or stained slides being washed).
Note on F20: Tips & Advice: • If your lab uses one or more type of specialised water, ensure lab users are informed which applications require which water types. • Soaking materials like dirty glassware is a more effective means of cleaning. • Relevant information should be in inductions.
Note on B22: This criteria should be completed as an alternative to research quality criteria, specifically for teaching laboratories.
Note on E22: Why? The greatest impact of a teaching lab will be the knowledge that the students walk away with. Ensuring that resource awareness and sustainability are integrated in teaching practicals can have a lasting effects beyond the course itself.
Note on F22: Tips & Advice: • Educate students about the carbon impact of the research they are currently conducting, or may conduct in the future. • Inform students on practical ways they can reduce the impact of future research. • Topics which may be pertinent include but are not limited to: fume cupboards and ventilation, chemical management, purchasing, waste segregation, health & safety interaction with sustainability, and equipment management.
Nature article on lab reducing lab plastics
UoEdinburgh freezer good practice guide
EPA guidance on small lab environmental practices
Article on single-use plastics in labs
Paper on impact of biomedical waste on city environment
Waste management guidance paper
Tufts example on take-back programs
Article on lab take-back schemes
KCL freezer management SOP
UoEdinburgh ULT freezer study
Efficiency exchange article on equipment sharing
Jisc article on equipment sharing
Cambridge global equipment sharing site
Paper on how room temperature affects lab equipment
UoColorado efforts to raise ULT temperatures
Long-term ULT freezer temperature study - UoEdinburgh
NIH guidance on lab keeping a lab notebook
Paper on publishing negative results
Paper on importance of publishing negative results
Nature article on importance of negative results
Paper on importance of core facilities for breakthroughs
Labs21 sizing guide for equipment
Princeton guidance on autoclave use
Article on keeping up lab books
NCSC article on backing up data
UoNewcastle guidance: core facilities acknowledgement
Article on planned maintenance for lab equipment
Cornell presentation on alarming via BMS
MIT chemical substitution portal
PDF of 12 principles of green chem. from ACS
NIH fact sheet on core facilities
UKRI guidance on fume cupboard maintenance
Harvard shut the sash programme
John Hopkins publication on fume hoods
HSE guidance on fume cupboards
Short video on closing sashes
Comparison of the 4 Types of Laboratory Water
Soak up water savings in the lab
Inter-university sustainable lab audits
Video on sustainability in the lab
Case study on KCL integrating sustainable practices
UoBristol's webpage on LEAF and how to take part

## 2. Bronze
Bronze
Total Criteria: 16
Category | # | Accreditation Level | Criteria | How did you meet this? | Useful Links
Waste | 1 | Bronze | The lab possesses required waste bins (possibly clinical, glass/sharps, hazardous etc.), as well as recycling/general waste bins with appropriate and clear signage. | Write how you've addressed this here, or why it's not applicable |  | P
People | 2 | Bronze | The lab has an induction procedure in place for all new arrivals, explaining the sustainable practices to take.
People | 3 | Bronze | The lab has a system in place to clear or track materials left by departing staff.
People | 4 | Bronze | Either the lab has a nominated person to drive sustainability forward or a group of people that meet to address sustainability within the lab.
People | 5 | Bronze | The lab (or relevant group) has taken part in 1 team activity of sorts over the course of the year, or one is imminently planned.
Purchasing | 6 | Bronze | Energy and materials consumption have been considered during the purchase of new equipment.
Equipment | 7 | Bronze | Heat sources on cold storage equipment are not blocked, and any filters are cleaned regularly.
Equipment | 8 | Bronze | Cold storage, ovens, or incubators are only operated when they are as full as possible.
Equipment | 9 | Bronze | There is a system in place to ensure equipment and lights are turned off when they are not needed.
IT | 10 | Bronze | Computer monitor brightness settings and computer time-to-sleep have both been minimised.
Sample & Chemical Management | 11 | Bronze | All samples and chemical containers have legible labels, or there is a system in place to ensure that going forward all samples will be consistently labelled.
Sample & Chemical Management | 12 | Bronze | The lab has a system in place for sharing chemicals between users within the lab group.
Research Quality | 13 | Bronze | Common protocols and methods are centrally shared and available to all lab members.
Research Quality | 14 | Bronze | The lab has had its pipettes calibrated in the past year, or has them scheduled to be done. In absence of pipettes, the lab has considered calibrating any materials commonly utilised for measurement.
Ventilation | 15 | Bronze | Any issues that estates must address have been reported. This includes ventilation, room pressure, water leaks, heating & cooling, and etc.
Ventilation | 16 | Bronze | Fume cupboards and safety cabinets possess signage encouraging good practice.
Case study on recycling plastics in tissue culture spaces at KCL
UCL permanent and daily departure documents may be found here
Article on considering efficiency to equipment
CU Green Lab Freezer guide including good filter practice
KCL policy on freezer good practice
Paper investigating signage effectiveness for recycling
UoEdinburgh case study on turning equipment off
UCL sticker templates to turn off lights and equipment
University of California ULT freezer tests
US dept. of energy case study on effectiveness of stickers on fume cupboards
Paper on fume cupboard savings in labs
UK government advice on reporting non-medical device failure
Paper on how standardised methods are important in TBM research
Article on reagent sharing
UoBoulder case study on benefits of sharing
Article on managing chemical inventories
UoIllinois chemical safety labelling advice
Article on common labelling errors in labs
Article on labelling methods
Paper on consequences on varying clinical methods
MIT guidance on scientific writing
ISO standards on pipettes and calibration
Paper on qPCR reliance on calibrated pipettes
Link to free UCL lab induction ppt.
Paper on importance of inductions for team effectiveness
Bringing recycling into the lab
UoYork article on lab plastic waste
Leading by example: Going green in the lab
How going green can raise cash for your lab
Effects of communication failures in operating theatres
How a group activity can increase social ties
The importance of team building and skill training
ULT Freezer guidance
Harvard guidance on reducing monitor brightness
Step-by-step guide to power management on different operating systems
University of Pittsburgh guidelines for departing investigators
UC Davis guidance on ways to save water and report issues
UCL Sustainable Equipment Guide, including purchase, operations, and metering
Note on E4: Criteria Below you will find all criteria you should address for Bronze. If you scroll your mouse over each criteria, background on why this criteria is important will appear as a comment.
Note on F4: How? Please write in how your lab has addressed each of the criteria below. You will find tips and advice as comments when you hover your mouse over each box.
Note on G4: Useful Links Click on links below which may facilitate implementing the criteria, including free resources.
Note on E5: Why? Laboratories produce immense amounts of waste (often plastics), much of which is incinerated at high-temperatures. Reducing the waste produced as well as treatment the waste receives can have significant environmental benefits, as well as reduce associated costs. Clinical waste will cost 3-5 times more than typical waste streams.
Note on F5: Tips & Advice: • Unless you work within a specialist laboratory (e.g. CL3), not all materials entering a laboratory are contaminated. Discuss with local H&S what is actually contaminated. • When introducing new waste streams, always ensure H&S are on board, estates have communicated a method of disposal and cleaners have been consulted, and users have been educated with training materials updated. • Consider the location and shape of bins for new waste streams. The volume of waste shouldn't vary, only the location/shape of bins.
Note on E6: Why? While there may be a substantial amount of information for new starters, it is crucial that efficient and sustainable practices are emphasised on arrival. It is during this period of initiation that people are most impressionable, despite there being a lot of new information.
Note on F6: Tips & Advice: • Consider the format of your inductions - is there a better way to reach someone than reading material? • Are there good practices which have yet to be formalised in your induction? Try and capture any ongoing good practice so that it may be continued. • Sustainable practices should not be separate, but rather fully integrated into induction materials.
Note on E7: Why? A primary cause of excess stored materials is due to the high turnover of laboratories, and the challenge in tracking what is left behind. Efficient tracking of materials will increase storage space as well as remove liability on those who must decide if something is to be thrown away. Tracking of samples and chemicals can ensure remaining staff have access to samples/chemicals/equipment which may have otherwise gone to waste.
Note on F7: Tips & Advice: • Keep your document or tracking method short and easy - often people departing will be under great pressure for completion. Make the process as simple as possible. • Try to make tracked samples (or etc.) accessible to all in the lab. • Consider tying their sign-off at the senior level with the tracking of their materials to lower risk of it being skipped.
Note on E8: Why? Engagement is most effective when driven by local examples of best practice rather than as a requirement. Nominating one or more people provides a reference for best practice and point of contact for those not able to fully engage. Ensuring continuity is crucial to this criteria. Succession planning is vital to ensure that good practice is not lost with departures and is truly engrained within the running of the lab.
Note on F8: Tips & Advice: • Green groups help raise engagement and can spread responsibilities, enabling more to be accomplished. Try to engage your colleagues on creating a ‘green lab group’ which meets 3-6 times per year. • Consider gaining senior support for having such a role specified. They may ask for validation, in which you may cite savings as estimated by this tool. • Including relevant responsibilities in a future Job Description is a possible way of ensuring continuity.
Note on E9: Why? Fostering relationships for new and existing staff and students will not only encourage collaboration, it will improve the overall environment for everyone.
Note on F9: Tips & Advice: • Try to ensure that as many members of your lab are present, and invite neighbours if possible. • Avoid activities which are solely dependent on alcohol. • Consider other activities and means to improve workplace wellbeing.
Note on E10: Why? If unit price is the prime consideration during purchasing, the lab runs a risk of costing more through increased energy/consumable/servicing costs. Furthermore manufacturers that are able to consider energy/material efficiency typically produce the best equipment as they have the resources to consider such areas. Thus purchasing the most sustainable equipment will often result in the best performing equipment and long-term cost savings.
Note on F10: Tips & Advice: • Prior to making a purchase check with your local estates/sustainability/energy team on any existing schemes which promote or incentivise efficient procurement. • Ask sales representatives to disclose all possible costs and requirements of new items. For e.g. has full servicing been factored in? Will estates need to take-over any maintenance once servicing ends? Can consumable remain fixed? How long is the response time for repairs, and what issues might that raise?
Note on E11: Why? ULT freezers in particular can produce immense amounts of excess heat. If blocked, this heat will drive the unit to be significantly less efficient. Excess heat can also cause room cooling to be triggered unnecessarily. Essentially any excess heating or cooling represents wasted energy.
Note on F11: Tips & Advice: • Consider if you ever feel a breeze or heating/cooling from a piece of equipment - is this blocked? • Consider if ULT freezers and similar equipment have heat blocked at the back by boxes or materials. • Never place a heat/cold source near a thermostat! This will cause the room's heating/cooling to become unbalanced.
Note on E12: Why? Most equipment that heats or cools will consume significant amounts of energy. Running half loads can use just as much energy as a full load. Thus running full load saves energy and prolongs the life of your equipment.
Note on F12: Tips & Advice: • Consider when users may require space in the machine, and always consult fellow staff that use the equipment in question. • Note that the larger items of kit will require more time to get back to temperature (warm or cold), consider recovery times if it is likely it will be needed again soon. A timer may assist in such situations.
Note on E13: Why? Within communal spaces, left on equipment and lighting is extremely common as it's sometimes difficult to know when something will next be used. Helpful signage is the first step in promoting good practice, and can prolong the life of equipment and lighting significantly.
Note on F13: Tips & Advice: • Consider if there is an existing standard for such signage available in your institution from your local sustainability/energy team. • Consult with all users prior to timer installations, and ensure they understand how to circumvent the timer if necessary. • Any equipment signage may require induction materials to be updated so that new members of staff are made aware of such systems.
Note on E14: Why? Reducing screen brightness and minimising the time until a computer sleeps are simple and effective methods to reduce energy consumption associated with IT work.
Note on F14: Tips & Advice: • Ensure screen brightness is not set so low as to cause eye damage. • Consult with local users about sleep settings, and consider engaging IT if they set-up your computers.
Note on E15: Why? Illegible samples are essentially waste, as they have little use if the contents are indiscernible (similar to chemicals in which labels have lost legibility). Instigating a simple and clear labelling system for samples will ensure you maximise the use of all materials present and avoid wastage.
Note on F15: Tips & Advice: • Ensure the writing utensils for labelling can endure your storage conditions (e.g. freezer and/or ethanol resistant markers). • Chemical labels from manufacturer are often sufficient, but consider indicating when they were received and expire, when they were opened, and ownership. If legibility is lost due to damage or time, redo the label. • Labelling systems should be joined up with any institute or standards required locally.
Note on E16: Why? Having communal chemicals is a simple way to save money, and can reduce wait times when missing a reagent. Some users may have valid concerns around contamination, in which case those selected chemicals should be removed from the sharing system. Try to consider the simplest method to promote sharing, and make the most of chemicals available - complicated systems may inhibit use.
Note on F16: Tips & Advice: • Local sharing may be via an excel, marked shelving, or a lab management software. • Ensure all staff understand the system in place, and also consider users' privacy and data protection. • Consider the upkeep of the system - how will it function in a few years time? How can it be the most useful while requiring the least amount of upkeep?
Note on E17: Why? While some variation in protocol methods may be necessary for research purposes, having standard methods within the laboratory will improve comparability between laboratories.
Note on F17: Tips & Advice: • Make the protocols editable if possible, but so that variations may be tracked. • Discuss appropriate protocols within lab meetings. • Protocols ideally should be aligned with any ISO standards, or comparable methods of labs doing similar research.
Note on E18: Why? Pipettes and other measurement devices are used extremely frequently in typical laboratories. Ensuring they are as accurate and comparable as possible can reduce error rates significantly, and avoid repeated experiments.
Note on F18: Tips & Advice: • Check with lab management if there is already a system in place. • Some companies can provide paid or free calibration. • Consider linking up with other labs/groups to do a larger calibration to minimise effort.
Note on E19: Why? Labs are often designed to maintain a particular negative pressure to ensure contaminants don't harm users. Not considering appropriate pressures or allowing excess heating/cooling may lead to significant amounts of energy being wasted, as well as compromised cleanliness. Any excess heating or cooling will waste energy, and typically wear down your equipment down quicker. Quick reporting of issues is important to ensure quick repair.
Note on F19: Tips & Advice: • Ensure you understand if a local fume cupboard for example requires a nearby door to be shut, and what pressures need to be maintained. These will likely have H&S implications. • Ensure users are aware of how to contact estates or support in case of an issue. • Inform users on any relevant heating or cooling policies for the lab, and where to go to should there be excess heating or cooling.
Note on E20: Why? Fume cupboards in particular often are responsible for the most energy consumption it the lab. You can mitigate this through sash closure, reduced flow rates, and putting into low-flow where applicable. Signage is a proven first step in promoting such good practice. Use signage to indicate how users may achieve simple good practice.
Note on F20: Tips & Advice • Note the difference between fume cupboards (protect users, limited sample protection) and safety cabinets (protect samples, as well as users). • There are 3 general types of fume cupboards. Constant volume, 2-speed, and variable-air volume (VAV). All may have flow rates potentially reduced (via containment testing by a professional). 2-speed should have sashes lowered and lower speed engaged when not in use. All units should have sashes lowered whenever possible. • Safety Cabinets should typically be off when not in use. Note local procedural differences for various uses of cabinets.

## 4. Gold
Gold
Total Criteria: 12
Category | # | Accreditation Level | Criteria | How? | Useful Links
Waste | 33 | Gold | There has been a recorded reduction of total waste produced, and/or an increase in recycling rates. | Write how you've addressed this here, or why it's not applicable |  | P
People | 34 | Gold | The lab has implemented at least one action to reduce travel.
Purchasing | 35 | Gold | LEDs have been considered for research illumination applications and purchased where feasible.
Equipment | 36 | Gold | There is a process in place for excess equipment and materials in the lab to be shared, repaired locally, or sold.
Equipment | 37 | Gold | Where water is used for cooling it is recirculated.
IT | 38 | Gold | Computing code has been optimised, and the number of storage clusters has been optimised to the tasks or schedule of tasks.
Sample & Chemical Management | 39 | Gold | Excess samples, materials, and chemical databases are made available to external laboratories, and/or the lab has made use of existing samples or chemicals from external sources, such as the UKCRC Tissue Directory.
Sample & Chemical Management | 40 | Gold | At least 75% of samples and chemicals are being actively used, or being stored and are easily identifiable. No more than 25% should be uncatalogued.
Research Quality * | 41 | Gold | The lab has adopted a laboratory management software, or has reviewed the options and provided a reason why this isn't appropriate.
Research Quality * | 42 | Gold | Sterilisation and cleanliness methods have been reviewed for efficiencies and effectiveness. Including but not limited to autoclave methods, UV sterilisation necessity, and cleaning rotas.
Ventilation | 43 | Gold | The lab has engaged and implemented actions via estates on lowering: fume cupboard flow rates, air change rates, and/or removing unnecessary extracts from safety cabinets to become recirculating.
Water | 44 | Gold | Guidance on appropriate usage of drains and effluent waste is communicated to all lab users, during inductions and beyond.
Teaching * | 41 | Gold | Environmental impacts are considered in the design or revision of experimental procedures for taught laboratory courses.
American chemical soc. guide to less lab waste
UoQueensland guide to less lab waste
E.g. of conference conducted online
Review of carbon footprint due to travel
Article on choice of green lab consumables
Article on managing lab consumables
Cambridge case on LED's & plant growth
Paper on solid state replacing argon lasers
Article on making use of old lab equipment
Article of recycling old lab equipment
Paper on reducing mercury in microscopy
UCL case studies, link kit repair workshop
US dept. of energy guidance on water pumps
UoManchester guidance on less lab water use
KCL case study on waterless condensers
Paper on optimised code reducing overheads
5 tips on enhancing code quality
Biobank UK website (UKCRC)
EU site for sharing bioinformatics
Sharing site for excess samples and materials
BBMRI-ERIC site for sharing EU samples
Paper on impact of LIMS on quality
NHS Wales case study on LIMS system
Paper on LIMS case study
KCL case study on reducing flow rates
CIBSE Guide B2 PDF
Environmental Sustainability in Biobanking
NIH guidance on decontamination
Paper on importance of cleanliness in teaching labs
UoCambridge guidance for discharges to drains
UoExeter guidance for discharges to drains
United States effluent guidelines
Article on challenges of sustainable biobanks
Sustainability in the teaching labs at UEA
MIT's Green Chemistry Wizard for alternatives
Vanderbilt's tips on sustainable teaching
Note on E4: Criteria Below you will find all criteria you should address for Silver. If you scroll your mouse over each criteria, background on why this criteria is important will appear as a comment.
Note on F4: How? Please write in how your lab has addressed each of the criteria below. You will find tips and advice as comments when you hover your mouse over each box.
Note on G4: Useful Links Click on links below which may facilitate implementing the criteria, including free resources.
Note on E5: Why? Initial bronze criteria focused on signage and training, while silver were intended to show correct segregation of waste i.e. training and signage is taking effect. This criteria now focuses on showing the impact of these efforts through improved recycling rates, or reduction of clinical/hazardous waste amounts.
Note on F5: Tips & Advice: • There are variety of ways to show improved recycling rates - LEAF contains a waste calculator which should be employed as evidence. • Consider the impact of any schemes the lab takes part in which reduce waste. • Consider your buying habits to truly reduce waste levels!
Note on E6: Why? The impact of extensive travel on the environment can equal or even surpass the impacts of laboratory research. It is recognised that travel may be necessary for the advancement of innovation, and thus this criteria does not seek to limit this. Rather the lab is encouraged to review its travel practices to identify opportunities which do not compromise output.
Note on F6: Tips & Advice: • Consider local travel between facilities as well as long-distance travel. • Engage with your facilities/administration on available telecommunications which may facilitate meetings. • Utilise your institution's travel service/agent, and request trips with the lowest carbon-impact where feasible.
Note on E7: Why? LEDs are becoming the go-to way of illumination for much microscopy, and solid-state lasers are a common option. LEDs don't heat samples as much, use a fraction of the energy, improve brightness consistency, can negate the need for toxic substances (mercury), and don't require warming-up. While they cost more initially, most units have been shown to save the lab money long-term. Furthermore LEDs may now be used for plant growth.
Note on F7: Tips & Advice: • Check with microscopy users if LEDs are already in place, and if there is a plan to replace should funds become available. • Ideally all old bulbs will be utilised prior to replacement of a unit, to make best use of existing stockpiles. • LEDs should be considered for plant-growth facilities.
Note on E8: Why? Old lab equipment is a common occupier of space within laboratories. Removing it can release space for other uses while becoming useful elsewhere. Repair and reuse of useful equipment is the goal of this criteria, while creation of space is a secondary benefit.
Note on F8: Tips & Advice: • Not all unused equipment should be treated the same depending on use, cost, ownership and etc.. Consider various routes for: equipment which may be disposed, which others may still have use for in their lab, or which may still be useful for parts. • Some services offer to collect and sell on old lab equipment. Consider these, as well as hiring local repair companies to repair equipment on-site. • Some repair companies will give discounts in exchange for parts - Engage with your local repair companies.
Note on E9: Why? Chilled water is a common method of providing cooling for reactions and equipment, and may be used to provide a vacuum. A variety of alternatives exist today though which can recirculate water or utilise mechanic pumps. Avoiding water going straight to drain can have the added benefit of reducing flood risk for unattended reactions.
Note on F9: Tips & Advice: • Consult local facilities and neighbouring labs if there are existing recirculating water chillers available. Note minimal use of water may not warrant purchase of a chiller. • A variety of waterless condensers exist which can remove the need for cooling in many condensation reactions. If applicable, the lab should implement these. • Small cheap pumps may supplement the need for water to facilitate a vacuum.
Note on E10: Why? Unoptimised computing code will increase time and energy required to perform operations. Similarly, excess clusters of server space will result in increased energy consumption required to achieve the same performance as optimal storage clusters.
Note on F10: Tips & Advice: • Consider the next user who may have to edit your code - will they understand your modifications? • Optimisation may come at cost of other functions/performances. Consider all implications of code optimisation. • Consider storage cluster numbers, not server numbers, when assessing cloud or external storage companies.
Note on E11: Why? Collection and storage of samples represents a huge investment in resources. If unused, this is essentially a waste. A growing number of initiatives are focused on making existing sample collections more widely available, through consolidation/publication of sample databases (with permissions). This criteria aims to make users consider if they've maximised existing sample databases, or considered interrogating existing databased rather than collecting new samples.
Note on F11: Tips & Advice: • Consider if there are existing materials available prior to collecting new samples. • Consider transport, location, and access rates of desired materials. Also consider any what kinds of sample data a biobank or facility may require. • All UKCRC tissue banks must fully register their tissue collections with the Centre and make their samples accessible for use by the wider research community.
Note on E12: Why? Labs commonly build up samples (e.g. in freezers) or chemicals and reagents, resulting in waste and unnecessary storage. Previous criteria targeted installing systems to ensure new chemicals and samples may be tracked. This criteria requires old obsolete samples and chemicals to be cleared out.
Note on F12: Tips & Advice: • Engage with relevant owners of materials. Explain rationale for requiring clear-outs. Communicate everything. • Consult H&S for assistance on chemical clear-outs. • A staggered approach may be necessary, as clear-outs can be time intensive.
Note on B13: This criteria is specifically for research labs. For teaching labs, the research quality criteria are not necessarily applicable. For research labs, please complete this criteria.
Note on E13: Why? Laboratory management software can facilitate a number of the best practices recommended by these criteria, as well as others. They can assist with H&S documentation, training, equipment booking/maintenance, purchasing, sample/chemical management, and more, while reducing the risk of data loss.
Note on F13: Tips & Advice: • Consider if there are existing software which you may access, at a departmental or institutional level. • When selecting a software, ensure you consider the following: all relevant activities and variances of users, long-term upkeep and maintenance of the software (financially and technically), benefits, time-requirement, and required training. • Software should be as user-friendly as possible to ensure maximum use. Avoid using many platforms requiring separate passwords.
Note on B14: This criteria is specifically for research labs. For teaching labs, the research quality criteria are not necessarily applicable. For research labs, please complete this criteria.
Note on E14: Why? Cleanliness and sterility are essential to most research practices. Achieving such cleanliness requires an immense investment (e.g. operating autoclaves, incineration of waste). There can be great variance in methods to achieve such cleanliness, some of which result in excess use of energy/water/plastics. Labs should assess if their methods to achieve a safe and sterile environment are sufficient, or possibly excessive.
Note on F14: Tips & Advice: • Consider the treatment of waste once is departs the lab - Is it necessary that current waste receives autoclaving? Is the autoclave operating at the right inactivation temperature/time? Are there any plastics which may be removed from the waste streams? • UV sterilisation for safety cabinets should be considered for efficacy. Is UV evidenced as effective for your work practices? • Are there clear responsibilities in terms of cleaning? E.g. do cleaners or staff know what may be cleaned, and who's responsibility it is if not cleaners?
Note on E15: Why? Extracting air represents a huge loss of energy for laboratories, as the air which is heated/cooled is then expelled to the outdoors (plus the operation of fans). As such, any reductions in extraction, when coupled with safe practices will save energy.
Note on F15: Tips & Advice: • Fume cupboard flow rates may be lowered to 0.5 m/s during annual inspection, and shouldn't rise greatly above (unless for specialist work e.g. radioactive). They may be lowered further if evidenced as safe via containment testing (typically a low cost) to 0.3-0.5 m/s depending on the model. • Air-change reductions must be done via estates, and the lab should only engage on feasibility. Note that recommended air change rates (the number of times air is changed in a room) for labs is quite broad (6-15 ACH according to CIBSE). • Ducted safety cabinets may be necessary, but sometimes they extract when not needed. Some cabinets may be altered without replacement. Replacements should only be made if the unit has reached the end of life.
Note on E16: Why? Putting harmful chemicals or materials down the sink may contaminate local water supplies or block systems.
Note on F16: Tips & Advice: • Ensure guidance is up-to-date with local H&S. • Users should be engaged beyond inductions, and signage should be ideally located visibly near drains. • Consider alternative solutions for materials which may block sinks but aren’t hazardous.
Note on B18: This criteria should be completed as an alternative to research quality criteria, specifically for teaching laboratories.
Note on E18: Why? Teaching labs can require large volumes of materials, consumables, and chemicals. Usage potentially may be reduced with consideration, particularly as teaching labs may not require the same levels of sterility as research laboratories.
Note on F18: Tips & Advice: • Consider applications of green chemistry - can any of the chemicals or reagents in use be replaced with less harmful or carbon intensive alternatives? • Consider what levels of sterility are necessary, and target reducing single-use items. • Consider ways to reduce the use of energy intensive equipment, e.g. reducing the needs for long-term sample storage, or reducing the need for fume cupboards.

## 5. Waste
Waste |  | Select 'Yes' when you want these figures to be included in the savings report | (select) |  |  | Progress Report
[6 empty] | Costs Saved (£) | -247.104 | -27.97%
Please fill in all yellow cells | [5 empty] | Carbon Saved (kgCO2) | -9.0688 | -16.67%
Select options in all green cells
Baseline Measure - Complete this section at the beginning of your study.
Waste Type | General | Recycling | Clinical/Hazardous | Offensive | Total
Number of Bins | 1 | 2 | 2 | 2 | 7
Collections per week | 2 | 1 | 3 | 1 | 7
Bin Bag Weight (approximate) | 0 - 3 kg | 4 - 6 kg | 4 - 6 kg | 0 - 3 kg | 0
Annual Cost | 24.596 | 55.77 | 772.2 | 30.888 | 883.454
Annual CO2 emissions (kg) | 4.5344 | 11.336 | 34.008 | 4.5344 | 54.4128
Progress Measure - Complete this section at the end of your study.
Waste Type | General | Recycling | Clinical/Hazardous | Offensive | Total
Number of Bins | 1 | 2 | 2 | 1 | 6
Collections per week | 2 | 2 | 2 | 1 | 7
Bin Bag Weight (approximate) | 4 - 6 kg | 0 - 3 kg | 4 - 6 kg | 0 - 3 kg | 0
Annual Cost | 61.49 | 44.616 | 514.8 | 15.444 | 636.35
Annual CO2 emissions (kg) | 11.336 | 9.0688 | 22.672 | 2.2672 | 45.344
Note on C10: General Waste This refers to any waste which is free of contamination and not being recycled.
Note on D10: Recycling Waste This refers to uncontaminated waste which is to be recycled.
Note on E10: Clinical/Hazardous This refers to any contaminated waste which is to be sent for incineration.
Note on F10: Offensive Waste Typically incinerated but at lower temperatures than clinical/hazardous waste, but which may be unpleasant to anyone who comes in contact with it.
Background
Although waste can be a challenging area to tackle, the potential savings in incinerated plastics/materials are great. Ensure that any variations of waste practices have been agreed to by local management.
Not all materials entering a Cat. Level II laboratory need to be incinerated.
Relevant Terms
Number of Bins: Please indicate the number of bins you have within your area for the relevant waste removal method.
Collections per week: Please indicate approximately how many times per week the relevant bin is emptied.
Bin Bag Weight (approximate): Please indicate approximately what the weight of the bag is. This is because most waste collection is charged per kilogram.
Tips
-You don't need to record every bag exiting your labs, but rather note the general rate for variation.
-If you install any supporting signage, ensure it is easily readable, and that users are comfortable with it.
-If uncertain, check, ask, and at times challenge. Policies may have changed, but always stay safe.
-Not all waste routes have been included, e.g. sharps.
Related Criteria
1, 17, 18, 33, 44

## 6. Fume Cupboards
Fume Cupboards |  | Select 'Yes' when you want these figures to be included in the savings report | (select) |  |  |  |  | Progress Report
[8 empty] | Costs Saved (£) | -854.8001 | -38.63%
Please fill in all yellow cells
Select options in all green cells | [7 empty] | Carbon Saved (tCO2) | -4.6259 | -63.4%
Baseline Measure - Complete this section at the beginning of your study.
|  | Fixed Air Volume |  |  | Variable Air Volume
| Type 1 | Type 2 | Type 3 | Type 1 | Type 2 | Type 3 | Total
Number of fume cupboards | 1 | [5 empty] | 1
Hrs On/Day | 24
Days On/Week | 7
Face velocity m/s | 0.6
Aperture (m2) | 0.6
% time sash open (VAV only) |  |  |  | 100%
Make-up air? | No | (select) | (select) | (select) | (select) | (select)
Region | London | (select) | (select) | (select) | (select) | (select)
Annual Cost | 2212.8841 | 0 | 0 | 0 | 0 | 0 | 2212.8841
Annual CO2 emissions (tonnes) | 7.2961 | 0 | 0 | 0 | 0 | 0 | 7.2961
Progress Measure - Complete this section at the end of your study.
|  | Fixed Air Volume |  |  | Variable Air Volume
| Type 1 | Type 2 | Type 3 | Type 1 | Type 2 | Type 3 | Total
Number of fume cupboards | 1 | [5 empty] | 1
Hrs On/Day | 24
Days On/Week | 7
Face velocity m/s | 0.5
Aperture (m2) | 0.6
% time sash open (VAV only)
Make-up air? | No | (select) | (select) | (select) | (select) | (select)
Region | (select) | (select) | (select) | (select) | (select) | (select)
Annual Cost | 1358.084 | 0 | 0 | 0 | 0 | 0 | 1358.084
Annual CO2 emissions (tonnes) | 2.6702 | 0 | 0 | 0 | 0 | 0 | 2.6702
Background
Fume cupboards typically use the most energy in a laboratory. This is because they require extract fans constantly running, which may have been heated or cooled to suit the laboratory. There are 2 distinct types of fume cupboards as shown left: Fixed Air Volume (or constant flow), or Variable Air Volume (or VAV). Fixed air volume hoods will maintain the same rate of extraction no matter the height of the sash, while VAV fume cupboards will modulate extract rates as the sash is lowered or raised. This modulation can result in significant energy savings.
How to tell if you have a Fixed or Variable Air Volume Fume Cupboard?
Most fume cupboards will have a small meter, either digital display or as a small speedometer-like meter with an arrow, which indicates approximate extraction rates. To tell the difference between a Fixed Air Volume and VAV fume cupboard, start with the sash completely open and check what the meter says for extraction rate. Now close the sash. If the meter displays a significant increase in extraction rate and it stays high, you have a fixed air volume fume cupboard. If the extraction rate modulates to stay within its original reading (+/- 20% max), then your cupboard is likely VAV.
Relevant Terms
Type: Three columns have been provided to permit you to enter a variety of types of fume cupboards, with the variation being in how they are used. E.g. you may have VAV units some of which are kept with shut sashes but some kept open, meaning you have 2 types of VAV fume cupboards.
Number of Cupboards: Indicate how many fume cupboards you possess of each type.
Hrs On/Day: For most research fume cupboards this figure will be 24 hrs/day. Only reduce this figure below 24 if fume cupboards are actually turned off completely (not just closing the sash), e.g. some teaching labs may turn fume cupboards off when not in use.
Days On/week: Similar to Hrs On/week, this should only be below 7 if fume cupboards are fully turned off.
Face Velocity: Each fume cupboard receives an annual inspection of extraction rates to ensure they are functioning. Engineers should indicate what their measured face velocity extraction rate is, which you can find typically on a small piece of paper on the cupboard. Please use this rate not what may be on the meter, they can differ.
Aperture (m2): For this, take the width of the fume cupboard (m) and multiply by the height (m) of the raised sash.
% time sash open (VAV only): Only for VAV fume cupboards - approximately what % of the time their sash is opened.
Make-up air?: Is there air which is entering the lab through an intended vent. You may need someone from estates to help determine this. If the air being provided is pulling on your doors, there's a chance you have no make-up air to compensate what is being extracted.
Region: This has been included to factor in that different regions will have to heat/cool the air more than others.
Tips
There are a few ways for users to reduce fume cupboard energy consumption which have an enormous impact:
1. Reduce face velocities to 0.5 m/s or lower - To reduce below 0.5 m/s containment testing must be conducted.
2. Ensure that sashes are kept closed as much as possible, though this will only reduce energy consumption in VAV models, it still is good H&S practice. Otherwise turn off units if possible.
3. Keep your fume cupboard clear of materials as this may impede air-flow.
Related Criteria
6, 15, 16, 23, 30, 31, 42

## 7. Safety Cabinets
Safety Cabinets |  | Select 'Yes' when you want these figures to be included in the savings report | (select) |  |  | Progress Report
[6 empty] | Costs Saved (£) | -1095.12 | -71.43%
Please fill in all yellow cells
Select options in all green cells | [5 empty] | Carbon Saved (tCO2) | -2.1532 | -71.43%
Baseline Measure - Complete this section at the beginning of your study.
| Type 1 | Type 2 | Type 3 | Type 4 | Total
Number of Cabinets | 6 |  |  |  | 6
Power demand (Watts) | 180 |  |  |  | 180
Hrs usage/day | 24 |  |  |  | 24
Days/week | 7 |  |  |  | 7
Ducted? | Yes | (select) | (select) | (select)
Annual Cost | 1533.168 | 0 | 0 | 0 | 1533.168
Annual CO2 emissions (tonnes) | 3.0144 | 0 | 0 | 0 | 3.0144
Progress Measure - Complete this section at the end of your study.
| Type 1 | Type 2 | Type 3 | Type 4 | Total
Number of Cabinets | 6 |  |  |  | 6
Power demand (Watts) | 180 |  |  |  | 180
Hrs usage/day | 8 |  |  |  | 8
Days/week | 6 |  |  |  | 6
Ducted? | Yes | (select) | (select) | (select)
Annual Cost | 438.048 | 0 | 0 | 0 | 438.048
Annual CO2 emissions (tonnes) | 0.86127 | 0 | 0 | 0 | 0.86127
Background
For estimation of energy consumption, please only assess class II safety cabinets.
Relevant Terms
Type: Four columns have been provided to permit you to enter a variety of types of safety cabinets, with the variation being in how they are used. E.g. some units may operate all day while some are kept off.
Number of Cabinets: Please indicate the number of class II safety-cabinets you have. Note various columns have been provided for variances in units
Power demand (Watts): This piece of information is typically found in the technical manual of your cabinet, or any supporting marketing material. Look for a wattage with a range of approximately 75-250 watts. Note that you can get an instantaneous accurate reading with a power meter, but contact Sustainability for support on this.
Hrs usage per day: Please indicate how many hrs per day the cabinet is left ON, not just how long it is in use.
Days/week: Indicate how many days per week the cabinet is in use.
Ducted?: Cabinets may or may not be ducted. To learn if it is, look at the top of the cabinet to see if there is a duct connecting the extract to the wall/ceiling. Your local lab manager should know also whether they are ducted, as unducted models will require regular filter changes.
Tips
1. Turn it off when not in use. A booking system may assist with this (either digital, or even written on the cabinet).
2. Cabinets have an in-flow and down-flow. Try and make sure your in-flow rates don't greatly exceed 0.4 m/s (target of no greater than 0.55 m/s) during annual inspection.
3. Consider the effectiveness of UV lights running after usage - Is there sufficient evidence for your work that this improves sterility? Can they be operated for less time?
Related Criteria
6, 7, 15, 16, 23, 31, 43

## 8. IT
I.T. |  | Select 'Yes' when you want these figures to be included in the savings report | (select) |  |  | Progress Report
[6 empty] | Costs Saved (£) | -12.6187 | -37.87%
Please fill in all yellow cells
Select options in all green cells | [5 empty] | Carbon Saved (tCO2) | -0.02481 | -37.87%
Baseline Measure - Complete this section at the beginning of your study.
| Type 1 | Type 2 | Type 3 | Type 4 | Total
Number of Computers | 1 |  |  |  | 1
Hrs usage/Day | 5 |  |  |  | 5
Days/Week | 7 |  |  |  | 7
Type of Computer | Desktop PC | Desktop PC | Desktop PC | Laptop
Number of Monitors | 1 | 1 | 1 | 1 | 4
Screen Brightness | Bright | Bright | Bright | Bright
Time until monitor sleeps | 5 mins | 5 mins | 15 mins | 15 mins
Basic or advanced use | Basic | Basic | Basic | Basic
Annual Cost | 33.3212 | 0 | 0 | 0 | 33.3212
Annual CO2 emissions (tonnes) | 0.065515 | 0 | 0 | 0 | 0.065515
Progress Measure - Complete this section at the end of your study.
| Type 1 | Type 2 | Type 3 | Type 4 | Total
Number of Computers | 1 |  |  |  | 1
Hrs usage/Day | 5 |  |  |  | 5
Days/Week | 7 |  |  |  | 7
Type of Computer | Laptop | Laptop | Laptop | Laptop
Number of Monitors | 1 | 1 | 1 | 1 | 4
Screen Brightness | Bright | Bright | Bright | Bright
Time until monitor sleeps | 15 mins | 15 mins | 15 mins | 15 mins
Basic or advanced use | Basic | Basic | Basic | Basic
Annual Cost | 20.7025 | 0 | 0 | 0 | 20.7025
Annual CO2 emissions (tonnes) | 0.040704 | 0 | 0 | 0 | 0.040704
Background
Computing and IT equipment is now found in essentially all laboratories, and can be responsible for significant consumption of energy. In recognition of this, we have included basic computing equipment. High performance computing is considered in IT criteria, but not quantified due to variation in how servers may be managed etc..
Relevant Terms
Type: Four columns have been provided to permit you to enter a variety of types of IT equipment, with the variation being in how they are used. E.g. Some computers may be operated all day while some are kept off.
Number of Computers: The number of computers you have, not the number of monitors. E.g. One laptop with 1 supplemental monitor is 1 computer, 2 monitors.
Hrs usage/Day: How many hours per day the computer is typically in active use (i.e. someone is using it, or it's displaying a video/in the midst of processing.
Days/Week: How many days per week the computer is in use typically.
Type of Computer: Whether the computers are laptops, or desktop units.
Number of Monitors: Number of monitors in total, including any laptop monitors.
Screen Brightness: There are 3 settings you may select from. Bright is for when the brightness is on maximum (75-100%), mid would equal 40-75%, and anything lower than 40% would be Low.
Time until monitor sleeps: The amount of time the computer is set to have the monitor on until the monitor is put to sleep. This may include time until the whole computer sleeps, but is not limited to this in recognition that some may be set to keep operating with the monitor off.
Basic or advanced use: Basic use refers to any standard computer use, e.g. word processing, internet browsing, emails, music, and etc.. Advanced use refers to when the computer is heavy use, e.g. when operating many programmes at the same time, or processing large amounts of data.
Tips
1. Lower brightness to a level that doesn't strain eye sight.
2. Operate computers only when necessary, and have monitors sleep when not in use.
3. Reduce the number of monitors if they aren't necessary.
4. For high performance computing, consider the PUE of servers.
Related Criteria 10, 25, 38

## 9. Cold Storage
Cold Storage |  |  | Select 'Yes' when you want these figures to be included in the savings report | (select)
ULT Freezers | *For background, definitions, and tips, scroll down |  |  |  | -20 Freezers |  |  |  |  | Fridges |  |  |  |  | Progress Report
Please fill in all yellow cells |  |  |  |  | Please fill in all yellow cells |  |  |  |  | Please fill in all yellow cells |  |  |  |  | Costs Saved (£) | -965.666 | -14.44%
Select options in all green cells |  |  |  |  | Select options in all green cells |  |  |  |  | Select options in all green cells
[15 empty] | Carbon Saved (tCO2) | -1.183 | -11.85%
Baseline Measure - Complete this section at the beginning of your study. |  |  |  |  | Baseline Measure - Complete this section at the beginning of your study. |  |  |  |  | Baseline Measure - Complete this section at the beginning of your study.
| Type 1 | Type 2 | Type 3 | Total |  | Type 1 | Type 2 | Type 3 | Total |  | Type 1 | Type 2 | Type 3 | Total
Number | 3 | 2 |  | 5 | Number | 3 | 1 |  | 4 | Number | 2 | 2 |  | 4
Age | 2 - 6 years | 2 - 6 years | (select) |  | Age | 2 - 6 years | 2 - 6 years | (select) |  | Age | 2 - 6 years | 2 - 6 years | (select)
Room is mechanically cooled? | Yes | Yes | (select) |  | Alarms? | Yes | Yes | (select) |  | Ambient temperature | 21 - 25 C | 21 - 25 C | (select)
Alarms? | Yes | No | (select) |  | Ambient temperature | < 21 C | 21 - 25 C | (select) |  | Doors Opened | Moderately | Infrequently | (select)
Ambient temperature | 21 - 25 C | 21 - 25 C | (select) |  | Operating temperature | -20 C | -22 C | (select)
Operating temperature | -80 C | -80 C | (select) |  | Doors Opened | Moderately | Infrequently | (select)
Filters Clean? | No | No | (select) |  | How Frosted? | Frosted | Seals Defrosted | (select)
Doors Opened | Moderately | Infrequently | (select)
How Frosted? | Frosted | Frosted | (select)
Annual Cost | 3143.775 | 1870.44 | 0 | 5014.215 | Annual Cost | 787.258 | 317.627 | 0 | 1104.885 | Annual Cost | 290.168 | 278.78 | 0 | 568.948
Annual CO2 emissions (tonnes) | 4.6181 | 2.9108 | 0 | 7.5288 | Annual CO2 emissions (tonnes) | 1.1979 | 0.41609 | 0 | 1.614 | Annual CO2 emissions (tonnes) | 0.43288 | 0.41049 | 0 | 0.84338
Progress Measure - Complete this section at the end of your study. |  |  |  |  | Progress Measure - Complete this section at the end of your study. |  |  |  |  | Progress Measure - Complete this section at the end of your study.
| Type 1 | Type 2 | Type 3 | Total |  | Type 1 | Type 2 | Type 3 | Total |  | Type 1 | Type 2 | Type 3 | Total
Number | 2 | 3 |  | 5 | Number | 3 | 2 |  | 5 | Number | 1 | 2 |  | 3
Age | 2 - 6 years | 2 - 6 years | (select) |  | Age | 2 - 6 years | 2 - 6 years | 2 - 6 years |  | Age | 2 - 6 years | 2 - 6 years | (select)
Room is mechanically cooled? | Yes | No | (select) |  | Alarms? | Yes | No | (select) |  | Ambient temperature | 21 - 25 C | 21 - 25 C | (select)
Alarms? | Yes | No | (select) |  | Ambient temperature | 21 - 25 C | 21 - 25 C | (select) |  | Doors Opened | Moderately | Infrequently | (select)
Ambient temperature | 21 - 25 C | 21 - 25 C | (select) |  | Operating temperature | -20 C | -22 C | (select)
Operating temperature | -75 C | -70 C | (select) |  | Doors Opened | Moderately | Infrequently | (select)
Filters Clean? | No | No | (select) |  | How Frosted? | Frosted | Seals Defrosted | (select)
Doors Opened | Moderately | Infrequently | (select)
How Frosted? | Frosted | Seals Defrosted | (select)
Annual Cost | 1936.418 | 2051.118 | 0 | 3987.536 | Annual Cost | 815.728 | 495.254 | 0 | 1310.982 | Annual Cost | 145.084 | 278.78 | 0 | 423.864
Annual CO2 emissions (tonnes) | 2.7652 | 3.325 | 0 | 6.0902 | Annual CO2 emissions (tonnes) | 1.2539 | 0.83218 | 0 | 2.0861 | Annual CO2 emissions (tonnes) | 0.21644 | 0.41049 | 0 | 0.62694
Background
Cold storage can require huge amounts of energy to operate. Above you will find 3 tools to help you estimate the energy consumed by normal fridges, -20C freezers, and ultra-low temperature (ULT) freezers. ULT Freezers in particular are extremely energy intensive, and one unit can consume more energy per year than the average UK household. Improved care and maintenance of cold storage not only saves energy, it will make your equipment last longer and keep samples safer.
Relevant Terms
Type: Three columns have been provided to permit you to enter a variety of types of freezers/fridges, with the variation being in how they are used. E.g. you may have some units freezers which are defrosted, but some that aren't.
Number: Please indicate the number of relevant units you have. Note various columns have been provided for variances in units.
Age: Please select the age range which most resembles the units under assessment
Mechanically Cooled?: Is the room where your ULT freezers are stored cooled with an air-conditioner? If so indicate Yes. If they are cooled passively with a fan, indicate No. This doesn't apply to -20C freezers or fridges.
Alarms?: Please indicate if your freezers are alarmed by secondary alarms (not just the built-in temperature alarm). Note that you can vary the cost on the 1. Assumptions page. This doesn't apply to fridges.
Ambient Temperature: What is the approximate range of temperatures for the room containing the ULT freezers?
Operating Temperature: Select the temperature which best matches what your units are operating. This does not apply to fridges.
Filters Clean?: Are the filters (typically located at the front and bottom of the freezer) clean or dusty? Applies only to ULT freezers.
Doors Opened: Please indicate approximately how often doors are opened and closed. Infrequently indicates no more than 1-5 times per week. Moderately can indicate 5-20 times per week, and Frequently will indicate 20+ times per week.
How Frosted?: In particular, are the seals where the freezer shuts frosted? Do they impede the unit from closing? Also consider internal frost build-up, which will reduce the amount of usable space. Does not apply to fridges.
Tips
1. Minimise the amount of time that doors stay open, as this will require freezer compressors to kick into action as well as expose samples to warmer air.
2. Raise the temperature where feasible, particularly back-up units which are empty.
3. Keep units defrosted with seals intact.
4. Keep units in rooms where they don't overheat. While air-conditioning will require energy, it may be offset through gains on freezers operating more efficiently.
5. Allow heat from units to escape, either from the back or the filters in the front.
6. Purchase efficient units.
7. Store only what you need, and throw-out old materials to avoid purchasing new units.
Related Criteria 3, 6, 7, 8, 11, 21, 24, 36, 40

## 10. Any-kit
Equipment | Select 'Yes' when you want these figures to be included in the savings report |  | (select) |  | Please fill in all yellow cells |  | Progress Report
[5 empty] | Select options in all green cells |  | Costs Saved (£) | -4867.2 | -42.86%
[7 empty] | Carbon Saved (tCO2) | -9.5697 | -42.86%
Baseline Measure - Complete this section at the beginning of your study.
kW rating of equipment |  | 10 |  |  |  | 10
Hours of operation |  | 24 |  |  |  | 24
Days of operation |  | 7 |  |  |  | 7
Annual Cost |  | 11356.8 | 0 | 0 | 0 | 11356.8
Annual CO2 emissions (tonnes) |  | 22.3292 | 0 | 0 | 0 | 22.3292
Progress Measure - Complete this section at the end of your study.
kW rating of equipment |  | 8 |  |  |  | 8
Hours of operation |  | 24 |  |  |  | 24
Days of operation |  | 5 |  |  |  | 5
Annual Cost |  | 6489.6 | 0 | 0 | 0 | 6489.6
Annual CO2 emissions (tonnes) |  | 12.7596 | 0 | 0 | 0 | 12.7596
Relevant Terms
kW rating of equipment: Every piece of equipment which uses energy will have a kW rating, typically provided in watts (W) or kilowatts (kW). If provided in watts, multiply by 1000 to convert to kW. You may find this information in a user manual, or sometimes on a tag on the equipment itself. The challenge may be with equipment that varies in rates of utilization. When there is variation, estimate the average.
Hours of operation: This figure should indicate the number of hours the equipment is in operation. E.g. you may only use the equipment for 8 hours per day, in which you would input ‘8’.
Days of operation: This figure indicates the days per week the equipment is in operation. E.g. you may only use water for 5 days a week, in which you would input ‘5’.
*Totals are provided as an annual sum.
Tips
1. You may be able to measure kW using a power meter. Ensure you have permission to do so, and seek guidance on how to utilise any meters safely prior to use.
2. Some equipment may have a digital display which remains on while the equipment itself is not in use. E.g. a centrifuge may have a small display, but not be actively spinning. Target equipment kW when it is in full use, as most digital displays will utilise negligible sums of energy.
3. Avoid metering any 3-phase power equipment unless with a professional electrician.

## 11. Water
Water | Select 'Yes' when you want these figures to be included in the savings report |  | (select) | Please fill in all yellow cells |  | Progress Report
|  |  |  | Select options in all green cells |  | Costs Saved (£) | -1539.2 | -78.72%
Baseline Measure - Complete this section at the beginning of your study.
Volume of water used (Litres per hour) | 100 | 50 |  |  | 150
Hours of operation | 24 | 8 |  |  | 32
Days of operation | 7 | 5 |  |  | 12
Annual Cost | 1747.2 | 208 | 0 |  | 1955.2
Progress Measure - Complete this section at the end of your study.
Volume of water used (Litres per hour) | 50 | 50 |  |  | 100
Hours of operation | 8 | 8 |  |  | 16
Days of operation | 5 | 5 |  |  | 10
Annual Cost | 208 | 208 | 0 |  | 416
Relevant Terms
Volume of water used (Litres per hour): This is the amount of water which has been utilised. Note that it is recorded in litres (L) and not m3 (meters cubed). If your water was volume was noted in m3 please divide by 1000 to convert to litres.
Hours of operation: Your water usage may vary in terms of hours per day. E.g. you may only use water for 8 hours per day, in which you would input ‘8’. If you’re trying to determine the carbon or cost of an amount of water which doesn’t require any units of time, simply input ‘1’ here.
Days of operation: Your water usage may vary in terms of days per week. E.g. you may only use water for 5 days a week, in which you would input ‘5’. If you’re trying to determine the carbon or cost of an amount of water which doesn’t require any units of time, simply input ‘1’ here.
*Totals are provided as an annual sum.
Tips
- If you’re estimating a rate of water usage from a tap, you may need to determine the litres/hour of your tap which can vary. Small to normal kitchen taps typically flow at 4-6 litres/minute, or 240-360 litres/hour.
- This calculator estimates CO2 and cost associated with normal tap water. Any purified or specialist waters will likely have a much greater CO2 and cost associated, though this varies on what is being done to the water.
- Typically, water companies will charge both to provide and remove water (though the drain) with drainage often costing more. This calculator assumes the water being both provided as well as draining.
Related Criteria
32, 37, 44

## 12. Open Initiative
Open Initiative
(please provide brief descriptions for each)
Initiative | Comments | Quantifiable Benefit
One ULT Freezer purchased | It was chosen by the energy dept. and performs well | We may have saved 650 on the purchase
We purchased 2 efficient safety cabinets | They were cheapest, and performed the task we need | Users enjoy them more, and we saved maybe 400 on purchase costs due to selection support
Purchased a recirculating chiller | Attached to stop pouring water through to cool all the time. | Saves we estimate 12 m^3 of water per year.
Repaired 2 centrifuges | 2 centrifuges were about to be purchased, but instead we had 2 old units repaired. | We estimate we saved 7,000 quid this year.
How it works: Open Initiatives is a place where you can record all the good practice implemented which may not get captured by calculators. We suggest recording instances for e.g. where you saved in procurement due to sustainable considerations. Further examples include water savings, time saving exercise, fewer consumables, or examples of how good practice benefitted your laboratory. Please provide context for any quantified figures (e.g. litres of water saved per year). Below are examples in red, delete as you progress.

## 13. Savings Report
Waste |  |  | Savings summary
Progress Report |  |  | Total Savings
Costs Saved (£) | 0 | 0% | Costs Saved | 0
Carbon Saved (tCO2) | 0 | 0% | Carbon Saved | 0 | tonnes CO2
Fume Cupboards
Progress Report
Costs Saved (£) | 0 | 0%
Carbon Saved (tCO2) | 0 | 0%
Safety Cabinets
Progress Report
Costs Saved (£) | 0 | 0%
Carbon Saved (tCO2) | 0 | 0%
I.T.
Progress Report
Costs Saved (£) | 0 | 0%
Carbon Saved (tCO2) | 0 | 0%
Cold Storage
Progress Report
Costs Saved (£) | 0 | 0%
Carbon Saved (tCO2) | 0 | 0%
Equipment
Progress Report
Costs Saved (£) | 0 | 0%
Carbon Saved (tCO2) | 0 | 0%
Water
Progress Report
Costs Saved (£) | 0 | 0%
Savings Summary
Here you will find a summary of all the progress and good work you've done.
You can see how much you've saved in each tab (except open initiatives), and how much you've saved in both carbon and pounds. For savings to appear on this page ensure that the relevant calculators have had "Yes" selected in the green drop-down box, which asks if it should be included.